{"text_format":"3","workbook":"xls","n_sheets":3,"sheets":["Sheet1","Sheet2","Sheet3"],"read":"whole","date_system":"1900","formulas":"cached_only"}

{"sheet":"Sheet1","index":0,"n_rows":14,"n_cols":7,"rows_kept":14,"flags":["formulas_cached_only"]}
["序号","申购科室","设备名称","数量\n（台套）","单价\n（万元）","总额\n（万元）","功能需求"]
[1,"第二手术部","手术床",1,30,30,"电动控制升降、倾斜、平移等功能，满足术中透视；可适配骨科牵引床。"]
[2,"第一手术部","电动手术床",2,9.8,19.6,"电动控制升降、倾斜、平移等功能，满足术中透视。"]
[3,"第一手术部","无影灯",3,5,15,"提供明亮、无影的光线；灯头可灵活调节位置和角度。"]
[4,"EICU","有创呼吸机",1,35,35,"食道压监测、通气模式齐全"]
[5,"重症医学科","有创呼吸机",2,60,120,"带有胸部电阻抗成像功能"]
[6,"重症医学科","呼吸机",1,35,35,"带有机械能监测功能"]
[7,"重症医学科","心肺复苏机",1,16,16,"自带电源电控"]
[8,"重症医学科","振动排痰机",1,20,20,"采用三维震动技术"]
[9,"口腔科门诊","口内X射线机",1,9.5,9.5,"1. 内置曝光剂量显示：实时显示每次曝光剂量。                      2. 球管参数：电压：60kV和70kV可选。焦点：≤0.7mm。管电流：4-7mA。                                                                           3. 安装与定位：壁挂式，伸展臂灵活。满足360度任意角度定位，臂长≥170cm。                                                                4. 手动曝光调整：可手动调节曝光设置值。"]
[10,"口腔科门诊","牙科动力装置（种植机）",1,9.5,9.5,"1. 显示屏：≥7英寸彩色屏，显示种植图案。                           2. 机头适配：支持多种转速比（1:1、1:4.2、1:5、16:1、20:1、27:1），标配20:1弯手机。                                                 3. 扭矩：5.5N·cm - 80N·cm，高精度马达，跳动≤0.02mm。                                                                                 4. 脚踏控制：水量、程序、正反转、转速调节。                         5. 扭矩监测：实时显示，记录峰值扭矩。"]
[11,"口腔科门诊","网电源供电骨组织手术设备",1,60,60,"口腔手术室使用，满足口腔颌面外手术要求；"]
[12,"介入手术室","IABP",1,115,115,"1、显示屏：≥10英寸，同步显示ECG、血压（收缩压/舒张压/平均压）、氦气容量。\n2、移动性：主机车架可分离，适配ICU/导管室转运。\n3、安全监测：三级报警系统（如氦气泄漏、触发信号丢失），球囊破裂时自动抽气防栓塞。"]
[13,"血液病科科研","血细胞分离机",1,78,78,"1、离心系统：转速≥5000rpm，环状离心带设计，泵精度误差≤±6%。\n2、操作安全：空气探测、压力监测、三级血液回输保护；抗凝剂智能管理（全血/抗凝比6:1–20:1）。\n3、移动性：大滚轮设计，支持床旁操作。"]

{"sheet":"Sheet2","index":1,"n_rows":0,"n_cols":0,"rows_kept":0,"flags":["formulas_cached_only"]}

{"sheet":"Sheet3","index":2,"n_rows":0,"n_cols":0,"rows_kept":0,"flags":["formulas_cached_only"]}